Load Cell Calibration

#### and Pendulum Test

George Lyons

Lafayette College

December 6, 2002

Abstract

An experiment was preformed to design, build, and calibrate a binocular load cell so that the forces due to an attached pendulum can be analyzed.  Using a strain box to measure the change in resistance of the strain gauges soldered onto the load cell, a calibration curve can be determined by subjecting the load cell to known forces.  This calibration curve is programmed into a HP VEE program so that the strain measured by the strain box can be calculated into a force.  The strain box is connected to the HP VEE program via a data acquisition board so that the force on the load cell is calculated in real time.  Using a 4.04 lb pendulum, the maximum and minimum forces on the load cell were observed to be 6.06 and 2.28 lbs respectively and the frequency was measured at 1.2 Hz.

Subject Headings: load cell: binocular - strain – frequency – pendulum – HP VEE

## Table of Contents

Table of Contents	2

Summary	3

Introduction	6

Apparatus and Test Procedure	8

Results	10

Discussion	13

Conclusions	15

References	15

Appendix	16

## Summary

For this experiment, a binocular load cell was initially designed using TransCalc.  A block of aluminum six inches of length, one and a half inches in height, and half an inch thick was used as a starting point.  Figure 1 shows the basic geometry of a binocular load cell.

Figure 1

The goal of this load cell is to have a maximum strain very close, but not exceeding 1500  when a force of 40 lbs is applied.  Using TransCalc, a strain of 1466  was determined using the values for each of the variables shown in Table 1.

| Bending - Binocular |
| --- |
| Applied Force, F | 40 lbf |
| Beam Parameters |
| Distance between hole CLs, L | 3.4 in |
| Radius, r | 0.3 in |
| Beam width | 0.5 in |
| Beam height | 1.5 in |
| Minimum thickness, t | 0.165 in |
| Modulus of Elasticity | 10e6 psi |
| Gage Parameters |
| Gage length | 0.125 in |
| Gage factor | 2.085 |
| Calculated Values |
| Recommended distance between gage centerlines, Z | 3.43 in |
| Nominal gage strain | 1466  |
| Strain variation | 7.6% |
| Span at applied force | 3.057 mV/V |

Table 1

After the load cell was properly cut using a CNC machine, four strain gauges and four soldering pads had to be glued on.  Each strain gauge has to be connected to another strain gauge in a certain way so that the change in resistance can easily be measured.  This change in resistance is measured by a device called a strain box, which has four colored posts for attaching a load cell.  Figure 2 shows how the load cell is correctly wired and also shows the color of each wire that attached to the strain box.  Before measuring any strain, the strain box must be calibrated by making sure that the “amp zero” is set to zero, the proper gauge factor is set (2.085 for this case), and that the output voltage off the BNC connection on the strain box is adjusted to 1.0(0.01 V when a force of 41 lbs is applied.

Figure 2

Once this is accomplished, the load cell can now be calibrated by subjecting it to numerous known forces and recording the strain measured from the strain box.  By graphing the force vs. output voltage and doing a linear regression, an equation for the calibration curve can be determined.  This equation will take the output voltage from the strain box and convert it to the amount of force that is applied to the load cell.

Now that the conversion from output voltage to force is determined, the HP VEE program can be written for the load cell when a pendulum is attached.  This program will display the force on the load cell on a graph versus time, display the frequency of the pendulum using a fast Fourier transform (FFT), and also record all of this data to the hard disk.

After the HP VEE program is written, load cell is now ready to analyze the forces of a pendulum.  A pendulum with a weight of 4.04 lbs and a length of 0.6223m was used.  The theoretical frequency was calculated to be 1.264 Hz and the frequency measured by the FFT is 1.25 Hz.  The forces on the pendulum were a maximum of 6.06 and a minimum of 2.28 lbs.  This amount of force makes sense since the average is close to 4 lbs and the pendulum swung slightly less than 160 degrees.

## Introduction

Load cells are important measuring devices to determine the amount of strain that a member undergoes when a force is applied.  These measuring devices can determine how close a member is to failing and therefore they are necessary in any situation to determine how close a member is to failure.  Many structures are designed to a certain factor of safety depending on the importance and threat of injury to a bystander.  After all of the designs are made and the structure is built, the material must be examined to make sure the desired factor of safety is actually achieved.  Using load cells, the amount of strain subjected to a member can be measured and then the actual factor of safety can be calculated to ensure that the structure was built to the correct factor of safety.

This paper will present the results of building and calibrating a binocular load cell to determine the strain and force resulting from an oscillating pendulum.  The load cell was initially designed with TransCalc and then cut out of an aluminum block using a CNC machine.  Figure 3 shows the basic geometry of a binocular load cell.

Figure 3

The goal of this load cell is to achieve a strain close to but not exceeding 1500  when the maximum force of 41 lbs is applied.  Table 2 shows the dimensions used to get a strain of 1466  when the load cell is subjected to the maximum force.

| Bending - Binocular |
| --- |
| Applied Force, F | 40 lbf |
| Beam Parameters |
| Distance between hole CLs, L | 3.4 in |
| Radius, r | 0.3 in |
| Beam width | 0.5 in |
| Beam height | 1.5 in |
| Minimum thickness, t | 0.165 in |
| Modulus of Elasticity | 10e6 psi |
| Gage Parameters |
| Gage length | 0.125 in |
| Gage factor | 2.085 |
| Calculated Values |
| Recommended distance between gage centerlines, Z | 3.43 in |
| Nominal gage strain | 1466  |
| Strain variation | 7.6% |
| Span at applied force | 3.057 mV/V |

Table 2

The load cell was calibrated by putting known forces on it in order to make a calibration curve.  This data was then used to determine the forces resulting from a pendulum on the load cell.  In addition, the frequency of the pendulum was analyzed using a fast Fourier transform and compared to a theoretical calculation of the frequency.

## Apparatus and Test Procedure

After the binocular load cell was designed and cut out via the CNC machine, the four strain gauges needed to be attached in order to measure the strain.  In order to connect all of the wires properly, four soldering pads also had to be attached.  See Figure 4 for a diagram of the setup.

Figure 4

Once all of the strain gauges are attached to the load cell and the correct wires soldered on, then the colored wires were attached to the strain box.  The colors on the diagram of the load cell match up with the colors of the posts on the strain box.  This box contains a very carefully constructed Wheatstone Bridge.  Before any strain can be measured, this box must be calibrated so that no strain is measured when there is no force on it.  To do this, make sure that the “amp zero” is set to zero, the proper gauge factor is set (2.085 for this case), and that the output voltage off the BNC connection on the strain box is adjusted to 1.0(0.01 V when a force of 41 lbs is applied.  The last step is done so that when the output voltage is connected to the HP VEE program, it is possible to use a gain of 10 and not exceed the capacity of the acquisition board.

After the strain box is calibrated, measure the strain and voltage output of a varying force from 1-41 lbs in increments of four pounds.  Note that when using the strain box, the actual strain is the measured must be divided by four.  This is because the strain box contains two Wheatstone Bridges and we are analyzing multiple strain gauges at the same time.

Next, set up the HP VEE program to display the waveform and fast Fourier transform for an oscillating pendulum.  For this to work, a conversion formula to turn the output voltage from the strain box into a force must be calculated from the calibration curve from the varying force.  In this case, the conversion formula is given in Figure 5.

Figure 5

The layout of this program is located in the Appendix.  After the program is complete, the pendulum was attached and the data recorded the waveform and FFT as it oscillated.

## Results

A graph of the data collected for the calibration of the load cell is shown in Figure 6.  This graph shows the varying force applied to the load cell on the x-axis with the strain on the left y-axis and the output voltage on the right y-axis.  Both of these graphs show that their relationship to the force is linear and this agrees with the basic formula of how strain is related to force.  The strain is measured to be 1216  when the full 40 lbs is applied.  This is 17.1% lower than the TransCalc value (1466  that was calculated in a previous lab.

Figure 6

The second part of the experiment tested the effect of a pendulum on the load cell.  Assuming the pendulum works as a simple pendulum (all of the mass is located as a point on the bottom), we can calculate a theoretical frequency for the oscillating pendulum by using the formula in Figure 7.  This value must then be multiplied by two because the HP VEE program records the pendulum swinging back and fourth as two cycles, whereas this formula calculates it as one.

Figure 7

It is interesting to note that this equation does not include the mass of the pendulum.  This is a result of assuming a simple pendulum.  Since no mass is distributed along the length of the pendulum, the amount of mass at the end does not affect the frequency and therefore the mass of the pendulum cancels out of the equation.

For this case, the pendulum has a length of 0.6223 m and the frequency is calculated to be 1.264 Hz.  Five hundred data points were recorded by the HP VEE program sampled at 50 Hz.  Figure 8 displays the change in force measured by HP VEE versus time.  We can tell that the pendulum is at its lowest point when the force is at a maximum, and the opposite is true when the pendulum is at its highest point.

Figure 8

By using a fast Fourier transform, the frequency spectrum can be displayed.  Figure 9 shows the FFT and the maximum frequency occurs between 1 and 1.5 Hz.

Figure 9

## Discussion

After all of the data was collected and analyzed, the theoretical and measured values can be compared to determine accuracy of this experiment.  First, it is important to examine how well this type of load cell performed in this experiment.  After all of the strain gauges were attached correctly, and the wires soldered on properly, this binocular load cell did not fail during any measurements.

The binocular load cell has a clear advantage over a simple cantilever beam that was used as a load cell in a previous lab.  In order for a cantilever beam to undergo the same level of strain, but keeping the width and length constant, the height must be greatly decreased.  This will make it harder to attach all of the proper wires and strain gauges.  Another disadvantage is that the cantilever beam might bend significantly since the height is reduced.  This could cause the material to swing slightly when performing the pendulum experiment, which then would cause errors in the data.  Since the binocular load cell performed well in all parts of the lab, it is suitable for this kind of experiment.

In the results section, the difference in strain measured and theoretical strain is 17.1%.  Neither of these values are incorrect and a careful analysis will show why.  The strain measured is determined by the strain gauges.  These gauges have a length (0.125 in) and the strain measured is the average strain over this distance.  The theoretical strain calculated by TransCalc determines the absolute maximum strain.  This is very important because the strain measured by strain gauges will not be able to determine the correct factor of safety.  In order for the correct factor of safety to be determined, a model of the load cell must be made using a force analysis program (such as ANSYS).  This program can determine the maximum strain, and also find the strain over the length of the strain gauge.  By averaging the strain over the length of the strain gauge, a theoretical value can be determined for what the actual strain gauge will read.  By using this method, the true factor of safety can be determined in the load cell.

The second part of this experiment analyzes an oscillating pendulum.  Before analyzing any data acquired, it is important to make sure enough data points were taken to avoid aliasing.  This occurs when the sample rate is too low and the correct frequency of the signal is lost.  According to Nyquist, the sample rate must be at least twice the frequency in order to record the correct frequency.  The theoretical frequency was calculated to be 1.264 Hz.  Therefore, the sampling rate must be at least 2.6 Hz to avoid aliasing.  Since a sample rate of 50 Hz was used, the data acquired accurately represents the frequency of the pendulum.

It is important to check how accurate the theoretical value of the frequency is to the measured value.  A total of 500 points were sampled at 50 Hz, which means that the program recorded data for 10 seconds.  During that period, the pendulum oscillated roughly 12 times, which means that the frequency is about 1.2 Hz.  This agrees with the theoretical value, but this estimate is not very accurate for analyzing the frequency.

Using a method called the fast Fourier transform (FFT), it is easy find the frequency of the pendulum with greater accuracy.  HP VEE has a built-in function to calculate the frequency spectrum of a signal and it can be displayed in a waveform (see Figure 9 in the appendix for the display of the HP VEE program).  Figure 8 in the results section shows the frequency spectrum of the pendulum by using a FFT and the maximum frequency was determined to be between 1 and 1.5 Hz.  Analyzing the raw data from the HP VEE, the maximum frequency occurs between 1.2 and 1.3 Hz.  This frequency confirms that the theoretical and the measured values are correct.  The theoretical calculation matched the measured value up to two significant digits, 1.2 Hz.  If a sample rate greater than 50 Hz was used, the frequency could be determined to more significant digits.

## Conclusions

The overall experiment was a success because all of the data obtained could be verified by the theoretical calculations.  The theoretical values for strain calculated by TransCalc and the frequency calculated for the pendulum were both very close to the measured values.  Once the average strain is calculated over the length of the strain gauge using a model, the theoretical strain could be calculated and this value would be very close to the measured strain.  All of the objectives for this experiment were answered with no unexpected results.

## References

Vishay Measurements Group, Inc. (1992).  Student Manual for Strain Gage Technology.  Pennsylvania: Vishay Measurements Group, Inc.

Wheeler, Anthony J., Ganji, Ahmad R. (1996).  Introduction to Engineering Experimentation.  New Jersey: Prentice Hall, Inc.

## Appendix

Figure 10 – TransCalc Data

Figure 11 – HP VEE program